## Slide 1: OP1.005 Telecon Talking Points
PZ Takacs
19 Feb 2014

## Slide 2: OP1.005 - Statistical Evaluation of Optical Surfaces
Presents standard methods for extracting statistical quantities relevant to optical surfaces from measurements of surface or wavefront topography.
Deals with the ISOTROPIC surface finish part of the surface topography, after the deterministic figure has been removed by LSF subtraction (DETRENDING). Applicable only to isotropic random rough surfaces. All others BEWARE.
  This DETRENDING is important. If it is not done correctly, the calculation will introduce artifacts into the results that will seriously distort the numbers.
  See OP1.004 for fitting polynomials to the figure part of the topography.
  The residual surface roughness must be able to be characterized by a random distribution, possibly with weak periodic features.
Presents specific formulae for computing functions in terms of discrete summations relevant to machine computation, rather than general definitions in terms of continuous integrals.
  We live in a digital world now.

## Slide 3: Motivation - Why this standard?
Modern interferometric instruments with digital cameras produce data arrays that contain much more information than just surface figure. Need a well-defined process for extracting high frequency information that doesn't introduce artifacts due to computational limitations.
Existing standards for surface topography (B46, ISO 4287 for profiles, ISO 25178 for areas) are designed primarily for machined surfaces that are characterized by a zoo of parameters. We only need to know the RMS roughness and PSD for optical surfaces.
B46 and ISO 4287 have not evolved very much from the old days of analog stylus profilers. Many outmoded concepts. Most are irrelevant for optically polished smooth surfaces.
  Discrete roughness and waviness filters. "Phase correct" filters to minimize distortion?
  Definitions are given by formal integrals rather than in discrete summation form.
  Frequency domain is not considered at all.
Make use of CPU computing power for Fourier analysis.
Allows for signal processing concepts in the frequency domain.
  Bandwidth-limited statistics by frequency filtering.
  Power spectral density (PSD) function analysis - direct connection to scattered light.
  Detrending, Windowing
  No need to transform back into height domain to compute statistics.
Old standards treat slope as a derived, or hybrid, quantity: derivative of height. Many instruments measure slope directly. Need to define slope as a separate quantity.

## Slide 4: Areal Slope definition (5.3.2)
Height is a scalar function of (x,y) position.
Slope is a vector function of (x,y). Described by projections onto x- and y-axes.
Earlier standards define profile slope in terms of profile height: [4287: 3.2.9]
But many instruments measure slope directly, so give it its own symbol: G for "Gradient" (5.3.2)
Areal Slope:
It has 2 components: Gx and Gy , each a function of x and y.
Linear profilers measure only one component at a time: Profile slope, Gx (5.2.2)
I only define areal slope. No discussion about fitting "slope surfaces" and computing residuals.
Note that the term "gradient" is used in 25178-2 [3.2.7] but is again defined as a hybrid parameter in terms of height.

## Slide 5: General comments
All the formulae contain the correct normalization factors to insure that Parseval is satisfied, i.e. power is conserved in the height and frequency domains.
Not addressed: zero-padding, standard data set with known results.
All Fourier transform expressions are given in terms of brute-force summations of the canonical forms of the expressions, where the index values are carefully defined. No off-by-one errors allowed.
In practice, actual computations will be done by canned functions and routines with hidden parameters relating to normalization and indexing. The brute force definitions provide a touchstone for insuring canned routines give the proper result.
In calculating statistics of random roughness profiles, Windowing distorts the appearance of a specific profile, but the ENSEMBLE AVERAGE over an infinite number of profile realizations is preserved.
  The Windowed profile may look funny, but its statistics are good.
We don't care what the profile looks like in the end, as long as the statistics are preserved over the ensemble average, because we stay in the frequency domain. The B46 people always transform back into the height domain, so appearance is extremely important to them.

## Slide 6: FT and PSD
The Fourier transform is defined to be the Discrete Fourier Transform, DFT (5.4.7)
The particular form of the DFT defined here puts the normalization factor on the inverse transform (5.4.8):
The PSD definition is in terms of the "periodogram" of the data set. (5.4.10)
PSD exhibits large fluctuations between points. There are standard ways to reduce the variance in the periodogram. Simplest is to average many to reduce the variance.
Smoothing of the periodogram are not addressed. Exercise for the reader.
Areal functions are analogous to the profile functions. (5.5)

## Slide 7: Band-limited statistics (5.6)
Detrending
Windowing
Old method - Transform, apply filter, inverse transform and compute statistics from height profile.
New method - Detrend and Window, then transform and compute PSD in frequency domain. STAY in frequency domain to compute statistics over well-defined bandwidth from PSD.

## Slide 8: Region of Interest, ROI
ROI is a necessary construct when dealing with circular aperture data on a rectangular camera, where data points are missing from outside the circle.
Need to define rectangular sub-apertures within the circular area that contain only "valid" data points. Avoid "bad" regions.
Bad regions can be NAN missing data, or outlier glitches. If there are only isolated bad points, possible to interpolate over them to maintain a large ROI.
  Interpolation methods are not addressed.
If more than one ROI is defined, each with a different number of points, difficulty arises in combining PSDs from each region due to different frequency scales. Can't just average the 2 together, unless the fundamental frequencies are the same.
  This is not addressed in the document

## Slide 9: Examples in Annex A
Examples of some of the concepts and functions defined in earlier sections are illustrated in Annex , (section 8).
Refer to the figures.

## Slide 10: Annex B
Several algorithms for height-to-slope and slope-to-height conversions are presented.

## Slide 11: Follow-on work (for someone else)
Surface description in terms of radial-azimuthal coordinates is not addressed.
Description of the PSD in terms of radial-azimuthal coordinates is not addressed. In many cases, it can be very useful to compute a radial PSD profile if the surface roughness is isotropic.